## Slide 1
Alberto Pullia
INFN - Milano University of Milano
Department of Physics
13th AGATA Week March 20-22, 2013 GANIL, Caen France
March 21-22, 2013
Status of the new AGATA digitizer*
*Technical details in white paper: DIGI-OPT12: 12-channel 14/16-bit 100/125-MS/s Digitizer with Optical Output for AGATA/GALILEO - version 1.7

## Slide 2
Mi-Pd group technical meetings from late 2010 to now
Oct. 6, 2010 - Legnaro
Dec. 15, 2010 - Padova
Febr. 1, 2011 - Legnaro
March 1 2011 - Padova
Apr. 21, 2011 - Padova
May 5, 2011 - Legnaro
June 21 2011 - Padova
July 19, 2011 - Legnaro
Oct. 13, 2011 - Legnaro
Dec. 21, 2011 - Legnaro
Jan. 17, 2012 - Legnaro
Apr. 27, 2012 - Legnaro
May 16, 2012 - Padova
July 12, 2012 - Legnaro
Sept. 18 2012 - Legnaro
Sept. 27 2012 - Legnaro
Oct 12 2012 - Legnaro
Dec. 11 2012 - Legnaro
Dec. 17 2012 - Legnaro
Feb. 5 2013 - Legnaro
12 chs digitizer card
Control card

## Slide 3
PS module, backplane
Valencia
System parts and connections (1 AGATA crystal)
Pre-processing Card
Padova
GTS
link
fibers
Clock-Distribution
and Control Unit
Padova
backplane
backplane
PS Module, backplane, cooling (under development)
48V
500 ps
Eye diagram @ 2Gb/s
backplane
MDR cables
3 Segment ADC Modules
(12 channels each)
Milano
Core ADC
Module
Milano
Workstation
PCIe expansion box

## Slide 4
ADC card
(this presentation)
Backplane for clock, control signals and PS
Backplane
connectors
Conceptual design
Use of a backplane greatly reduces cable burden
MDR connectors
Real thing
MDR connectors
Low noise
Low power consumption
High integration
High flexibility
ADC chips with integrated JESD204A encoder/serializer
FPGA-less design!
Key words:

## Slide 5
The card – top view
EMI filters
clock distr
analog front end
ADC, CK, TX manual reset
160 mm
120 mm
Power (3.3V, 2.0V) via backplane
ck, sync, test i2c, spi via backplane
MDR in
Power (3.3V, 2.0V) via screw header
SNAP12 TX (not connected)
conn’s for routing card
J5
J6
J7
ADC dual
quad digi-pot
Optional fast out with LLD
LEDs
ck, sync, test i2c, spi via mini-HDMI

## Slide 6
Time calibration by Damiano/Diego algorithm
+
Offset
Regulation
Fast
Amplifier
ADC
Antialias
Laser
Divider
Low
Low
I2C
SPI
Detector signal
The time calibration feature has been optimized and fully qualified @ LNL by Diego Barrientos
Damiano & Diego algorithm
I2C I/O
Ch by ch latency distribution due to JESD204A
Sync test signal

## Slide 7
Map of DIGI-OPT12 devices
SPI control of ADCs settings
I2C control of clock distribution settings
I2C analog MUX for time-calibration test pulse
I2C GPIO for range setting (7 and 21 MeV)
I2C digipots for offset setting (on a ch by ch basis)
I2C temperature meter
I2C volt-meter for PS check
I2C elapsed time meter (option)
and more …..

## Slide 8
Setup for testing the DIGI-OPT12 card
PC scripts
PS
DIGI-OPT12
Clock
Test signal
Notes: 8

## Slide 9
Setup for testing the DIGI-OPT12 card
Differential analog test signal Tektronix AFG3022B
Differential clock Tektronix AFG3252
I2C/SPI protocol interface Arduino Uno + custom shield
DIGI-OPT12 card
Laptop w USB link to Arduino
I2C/SPI
USB

## Slide 10
Setup for testing the DIGI-OPT12 card
USB link
Arduino UNO
Sony VAIO VPCZ1
Custom shield
I2C / SPI protocol three wires
Exchange data with Arduino through USB serial connection
Arduino IDE
Processing IDE
Communication through files in RAMDisk
DIGIOPT-12
Custom matlab scripts for real-time data analysis
Mini
HDMI

## Slide 11
Matlab GUI for digitizer testing
Signal in channel 1
The ADC chip has an undocumented 8 kB RAM where the waveforms can be stored and read out via SPI, which is used in the demoboard. I gained full control on that by reverse engineering. Very useful functionality for test benching and diagnostic.

## Slide 12
Clock signals captured at the ADC’s input pins
Clock signal @ 100 MHz is clean and symmetric !!
The clock signals are provided to the ADCs through the on board clock-distribution chip
Notes: 12

## Slide 13
Eye diagrams of the 12 high-frequency digital signals
Data stream = 2 Gbps per lane. Eye diagrams are nicely open !!
500 ps
The 12 encoded/serialized datastreams are provided to the optical TX at high frequency
Notes: 13

## Slide 14
Measurements - noise
7 MeV range
20 MeV range
Identical to the value rated in ADC datasheet: ENOB = 11.6 @ 70MHz !!
20 MeV range
Acquired data - no input signal provided

## Slide 15
Measurements – bandwidth and pulse shape
Anti-aliasing filter is set for a 26ns risetime in step-response
Exponential decay signal (1 MeV equivalent)

## Slide 16
Measurements with cross-connected demoboard
FFT (Hanning window, 20 MeV range selected)
FFT (Hanning window, 7 MeV range selected)
(*)
By putting sn = 1.48 or 2.2 in Eq. (*) :
SNR20MeV = 71.85 dBFS
SNR7MeV = 68.41 dBFS
These values are in good agreement with those provided by the demoboard software
Rated SNR = 72 dBFS
Rated SNR = 69 dBFS
DIGI-OPT12
Demoboard
3-wire SPI x-connection
Noise spectrum
Noise spectrum

## Slide 17
Card temperature with no heat sink on ADCs
USB-powered fan underneath (AEOLUS notebook cooler)

## Slide 18
Card temperature with heat sinks
Fan-less
With USB powered fan for notebooks
Fischer tiny heat sink
CK
IN
I2c spi
2.0V, 3.3V
PLL lock LEDs
Notes: 18

## Slide 19
First spectroscopy measurements with detector MARS
The noise seen on the preamplifier waveform is
231 uV r.m.s. (detector connected)
as compared with
180 uV r.m.s. (detector non connected)

## Slide 20
First spectroscopy measurements with detector MARS
After the trapezoidal filter the noise decreases to
0.50 LSB r.m.s in 16-bit format
from
1.89 LSB r.m.s. in 14-bit format

## Slide 21
Spectrum from small front electrode of MARS detector
We are now waiting for an AGATA detector crystal, which is needed for full qualification (should have been delivered to Legnaro in late October 2012 … )

## Slide 22
Specifications
ADC | 14 or 16* bit, 100MS/s with JESD204A (8b/10b) and SER interface
Channels | 12 per card
Analog input | Differential, MDR connectors as specified in AGATA preamp white paper v. 3.2
Configurations | “Segment mode” (default, 12 chs) or “Core mode” by insertion of passive piggyback PCB (3 chs + spares)
Clock input | Differential LVPECL through mini-HDMI or backplane
Sync and test pattern input | Single-ended LVCMOS with static toggle through e-SATA connector
Control input | I2C and 3-wire SPI through mini-HDMI (HDMI type C) connector
Output | Optical, 1 fiber each digitized channel (see datasheet of ReflexPhotonics SN-T12-C00601 SNAP12)
Power Supply | 3.3V @ 2.5A and 2.0V @ 0.6A
Power cons | ~ 10W per card
Size of card | 120mm x 160mm
Range control | Remotely controlled range selection: (a) “20 MeV range” (with 25% offset displacement) and (b) “7 MeV range”
Offset control | Remotely controlled within +-30% of full swing
ADC param control | Remotely controlled full set of ADC and JESD204A parameters (see datasheet of NXP ADC1413D)
Clock param control | Remotely controlled full set of clock-distribution parameters, includind switching on the embedded PLL for zero-delay option (see datasheet of AD9522-3)
Pulser param control | Remotely controlled full set of pulser parameters in “Core mode” (see AGATA preamp white paper v. 3.2)
Options | Interleaved mode (e.g. 6 equivalent channels @ 200MS/s), single-ended analog inputs, built-in clock generation
See also white paper: “DIGI-OPT12: 12-channel 14/16-bit 100/125-MS/s Digitizer with Optical Output for AGATA/GALILEO” version 1.7 (or later)
*Pin-to-pin compatible ADC, mod. ADC1613D, is available with a maximum sampling frequency of 125 MHz

## Slide 23
New pin-to-pin compatible ADC available
In 2012 IDT has commercialized a new ADC chip: ADC1443D
(1.8V, 3.0V) PS  single 1.8V PS with 44% less power consumption (!)
Pin-to-pin compatible to ADC1413D
Integrated JESD204B  deterministic latency
Total power of DIGI-OPT12 card:
presently: 9.7 W
with new ADC: 6.8 W
Using the new ADC1443D the power would go down to 0.57W/ch including the laser !!!
*ADC chips designed and made in France for IDT (previously NXP)
*
*
Original
New

## Slide 24
Road map
 April - June 2011
ADC1413D semi-qualified @ Padova
Damiano algorithm for cancellation of random latencies of ADC/state machine
 June 2011
Schematic diagram completed (in Orcad Capture)
 July - September 2011
Translation of schematics in other CADs (Zuken Cadstar)
 September - December 2011
Layout synthesis
 Spring 2012
First prototypes ready for testing
 Spring-summer-autumn 2012
Tests, qualification, preproduction for GALILEO
Early 2013
Production for 10 crystals of AGATA - running
Tests with an AGATA crystal - waiting for crystal